# Project Proposal

Group 14 - Diya Chakraborty, Hanna Ginther, Jenisha Thakar, Leila Tawfik, Emmanuella Osifo

February 23, 2026

BMEN 668

EXECUTIVE SUMMARY

The submitted project proposal presents a solution to reduce pain related to dysmenorrhea and chronic pelvic pain (CPP) caused by endometriosis. A vast majority of women diagnosed with endometriosis experience severe pain that can significantly reduce quality of life. Current solutions are either pharmacological (NSAIDs) or devices that are bulky and non-discrete. This project presents a low-profile device that takes a multimodal approach, targeting both central and peripheral pain associated with dysmenorrhea.

Based on currently available solutions, the following needs have been identified in potential users. Users require a device that addresses pain through psychological and physiological interventions rather than pharmaceutical approaches. These interventions should not cause harm to the user and should include automatic shut-off mechanisms. Additionally, there is a need to validate and assess user pain levels and frequency, which can be tracked through an app. Lastly, users require a device that can be worn under clothing in public without causing discomfort or embarrassment.

Although there is a gap in the market for the proposed solution, current solutions for endometriosis-related pain contain aspects that are valuable to build upon. There are devices that incorporate multimodal therapeutic elements; however, they lack the psychological component that is proposed. Another solution is discrete and portable, but it relies on hydrogel adhesive that is usable only for a finite period. The concept of a feedback loop to drive safe use of the device is also present in several existing designs. Overall, various current designs offer components that align with the project goal and can be adapted.

The potential design direction is broken down into several modules. The first is the hardware, which contains vibration motors and a heating pad within a discrete, flexible belt. Secondly, the software consists of an application that controls the device and is initiated by the user. The app can record pain levels and includes an interface through which the user can engage in relaxation exercises. Thirdly, the entire device is driven by a control loop. The hardware and relaxation exercises turn on once the user initiates the app, and the hardware has an automated shut-off mechanism once the relaxation exercises are concluded. A flowchart summarizing the user experience is included.

Technical specifications for the proposed device focus on user safety and wearability. Aspects such as thermal injury prevention, patient data security, and reliable power delivery are addressed, and target values as well as tolerances are provided.

Multiple purchases are required for the proposed project. The main hardware components include a microcontroller, vibration motors with a motor driver, a heating element, and a power supply. The application will be developed using free software, requiring no additional purchases. A summary of all required components is included.

A detailed risk analysis highlights technical, safety, and team performance risks. The highest risks are associated with hardware design and integration due to limited experience and testing opportunities. Mitigation strategies are outlined and primarily focus on maintaining technical specifications and clearly assigning roles and responsibilities.

TABLE OF CONTENTS

1.	Project Statement	1

2.	User Needs and Context	1

3.	Prior Art	1

4.	Potential Design Directions	1

5.	Technical Specifications	1

6.	Budget	1

7.	Timeline	1

8.	Project Risk Analysis	1

9.	Technical Specifications	1

Appendix A:	Additional Material	2

# Project Statement

Endometriosis is a chronic gynecological inflammatory condition, affecting approximately 10% of the female population of reproductive age, representing nearly 2 million Canadians [1].

Secondary dysmenorrhea and chronic pelvic pain (CPP) are hallmark symptoms of endometriosis, with approximately 90% of patients reporting severe or debilitating pain [2], [3]. These symptoms substantially reduce quality of life and disrupt daily functioning, with adolescents being ten times more likely to miss school and adults losing an average of nearly eleven hours of work per week [4]. Additionally, endometriosis-associated dysmenorrhea and CPP place a significant burden on the Canadian healthcare system, accounting for approximately 8,000 annual surgical and inpatient admissions and an estimated cost of nearly $25 million per year [5].

Medical therapy for these endometriosis-associated symptoms generally involves a combination of non-steroidal anti-inflammatory medications (NSAIDs) with oral contraceptives or gonadotropin-releasing hormone (GnRH) agonists [6], [7]. Non-pharmacological strategies include Transcutaneous Electrical Nerve Stimulation (TENS) and focal muscle vibration (FMV), mobile applications, virtual reality (VR), and wearable technology [8]. Specifically, FMV has been shown to significantly reduce pain scores in patients with  dysmenorrhea, while low-level continuous heat therapy demonstrates efficacy comparable to over-the-counter analgesics [9], [10], [11]. However, medical interventions often fail to provide adequate relief and are associated with adverse effects, including bone density loss, and infertility during treatment [6], [7]. Despite their potential, technological solutions face barriers preventing their widespread adoption [8]. Current devices are often bulky and non-discrete, which discourages use in public or professional settings. Furthermore, existing vibrational and heat therapies often rely on manual timers or lack automated shutoff mechanisms, impeding usability and posing safety risks such as “toasted skin syndrome” associated with chronic, unmonitored heat therapy use [8], [10]. Critically, there is a lack of accessible, responsive solutions that address both the physiological symptoms and the psychological aspects of chronic pain, such as the fear-avoidance cycle, where the anticipated threat of pain leads to inactivity and worsened physical outcomes [12]. Existing mobile tools remain passive logging platforms that fail to provide real-time therapeutic intervention of clinical insight.

This project aims to explore the feasibility of a low-cost, non-pharmacological approach for managing pain-related symptoms of endometriosis. The design will consider specific technical barriers, including the translation of pain-modulation concepts into functional features, the accommodation of high inter-user pain response variability, and the strict adherence to thermal and electrical safety limits. Success will be measured by the system’s functional reliability and its overall wearability. Our goal is to deliver a functional prototype that demonstrates the feasibility of a responsive system for the integrated management of endometriosis-associated pain, developed by a five-person team by April 2026, within a $170 budget.

# User Needs and Context

The primary end-users of this system are patients living with endometriosis who require improved pain management strategies. A central need identified is a physiological and psychological pain intervention strategy as an alternative to pharmaceutical treatment. Only 25% of patients experience relief from NSAIDs and these pharmaceutical interventions can cause infertility during treatment, indicating that multimodal interventions are often required to provide individuals with safe and effective options [6]. Therefore, an accessible non-pharmacological option that combines physical and psychologically based interventions can provide women with more ease and control with their pain-management options. From a physical perspective, users require heat and vibration therapy, along with guidance through pelvic floor-directed exercises to mitigate their pain symptoms [11]. Potential exercises include pelvic tilts, cat/cow, and child’s pose in addition to diaphragmatic breathing which has been shown to assist in pain management [13]. The psychological component of pain management is equally important in pain management with minimal intervention. The fear-avoidance cycle associated with chronic pain provides a useful framework where an anticipated threat of pain can lead to constant vigilance and monitoring, which may make even low-intensity pain feel worse [12]. Additionally, fear of provoking pain can result in avoidance of movement or activity, leading to inactivity and potentially worsening symptoms. Users therefore require structured breathing and guided movement strategies that actively intervene in this cycle to mitigate the perception of pain.

There is also a need to validate patient pain and symptom experiences, particularly among women with endometriosis. Improved symptom tracking and data collection can support validation by employing effective data storage solutions in addition privacy protections. A potential in-app community forum may further allow users to discuss and share experiences to support and validate their symptoms, depending on feasibility of development. Remote monitoring of symptoms represents an additional need for both patients and healthcare professionals. Existing belts provide no data collection, while symptom-tracking platforms provide no active pain management. When patients log and rank pain, the data can enable healthcare professionals to better understand the frequency and intensity of symptoms. It also allows users to gain a clearer understanding of when pain occurs and whether the intervention is effective. Longitudinal records of pain address validation need and facilitate proactive management of future flare-ups, shifting treatment from reactive to proactive strategies. Structured datasets could eventually support Machine-Learning-Model predictive approaches that could provide users with predictions on the timing and severity of their cycle as a future direction.

Another key need is a bio-responsive control loop to improve patient safety compared to “set-it-and-forget-it” heating pads. In this model, the patient logs and ranks their pain, the device activates, and the user is guided through exercises. If pain persists after the exercise phase, the cycle repeats. Therapy should only be active when clinically necessary (i.e., when pain is still present). Automatic shutoff mechanisms are required to reduce the risk of injury associated with prolonged heat exposure, including toasted skin syndrome [11]. This approach reduces overreliance on heat therapy and improves overall safety. It is also important to consider that endometriosis-related pain is unpredictable and does not occur exclusively at home. Many existing heating belts are large, clunky, or require connection to a power outlet which is not feasible when at work or on-the-go. As a result, the device must be small, discrete, transportable, and suitable for long-term wear. A low-profile, flexible belt with breathable fabrics and flexible electronics would allow pain management to function in various environments outside the home.

Secondary stakeholders include both patients and professionals who would be essential in validating and refining the system. Engagement with women experiencing endometriosis would be necessary to assess comfort, discretion, usability, and perceived effectiveness of both the wearable belt and the mobile application. Clinicians specializing in women’s health would be consulted to ensure that the guided exercises, safety shutoff mechanisms, and collected symptom metrics are clinically appropriate and meaningful. Women’s health app developers would be engaged to evaluate feasibility of implementation, data architecture, and user interface design. Biomedical engineers or product designers could provide input on hardware integration and wearability constraints. Additionally, consultation with data privacy or health informatics advisors would be important given the collection and potential remote sharing of longitudinal symptom data. Continued engagement with our professor will be necessary to monitor scope, ensure feasibility within project constraints, and maintain alignment with academic expectations.

# Prior Art

| Current Solutions (Product/Patent Name & Source) | Key Features (Novel Claims) | Strengths | Limitations | Proposed Design Direction Takeaways |
| --- | --- | --- | --- | --- |
| System and Device of Thermal Treatment with Protocol Based on Biological Feedback (RU 2565099 C2) | - System configured with a heat-exchange device to measure and quantify thermal energy transferred between the treatment pad and the human body. [14] - Uses measurable physiological sensors (e.g., heart rate, blood pressure, skin properties) to dynamically adapt temperature, pressure, duration, and frequency of treatment. Stores and analyzes biometric data over time to adjust treatment protocols [14]. | - Closed feedback loop for thermal therapy. - Individualized protocol based on measurable physiological data that is adaptive over time. | - Requires measurable physiological parameters and accurate quantification of heat exchange between the body and device. - Depends on biometric data collection and processing to regulate therapy. - Costly equipment and tedious set – up. | - This design has desirable features that would be beneficial to incorporate into our prototype, mainly the bio-feedback component loop. - However, measuring and quantifying heat exchange is not feasible within this semester. - The mechanical organization of this device is functional with a therapeutic treatment pad incorporated with biometric sensor(s) attached to a controller that analyzes input data and processes output treatment. |
| Cold/heat-assisted distributed wave vibration therapy (US10123937B2) | - Uses a grid/array of vibration motors on a medium that attaches to the body to create a traveling “wave” vibration sensation [15]. - Includes controllers to set vibration amplitude/frequency and generate wave vibration patterns with timing/decay between waves [15]. | - Multi-modal therapy (vibration + optional heat/cold) in one device. - Flexible/adaptable form factor that can conform to different body parts. - Customizable patterns (direction of waves, timing, intensity) enabling personalization. | - The “feedback” described is mainly user-controlled pattern creation (touch/shake/app input), not physiological biofeedback. - Designed broadly for general pain management/therapy, not specifically optimized for lower abdominal dysmenorrhea pain. - Does not inherently include integrated temperature sensing or automatic heat regulation to prevent overheating during prolonged use. | - Keep the idea of combined modalities (e.g., heat + vibration) and programmable therapy patterns, but upgrade to a safer, abdomen-focused wearable with integrated temperature monitoring and adaptive heat control. - Design a lower abdomen focused wearable with safety temperature sensing/limits and adaptive dosing to reduce overheating risk during prolonged use. |
| Portable combined stimulation device for alleviating menstrual pain (US9421123B2) | - A portable combined stimulation device which alleviates menstrual pain through TENS and a heater pad . - It is a pad which can be applied on the abdomen or back [16] | - Can be applied anywhere on the body to target pain - Heater pad works through a feedback loop to prevent heat related injuries - Has a rechargeable battery | - Hydrogel adhesion is reusable for a finite period of time. With extended exposure to environmental factors the adhesion can wear away. | - Heating pad feedback loop can be integrated into the design to prevent injury. The design is also slim and can fit under clothing for a discrete solution, this can be adapted for our specific design idea. |
| Wearable device to provide simultaneous relief from menstrual pain and cramp discomfort using dual frequency (WO2023053148A1) | Utilizes a dual-frequency approach that automatically and alternatively switches between two frequency ranges: 50-120 Hz and 1-4 Hz at present intervals [17]. Designed as a portable, single-unit device that incorporates a controller and microprocessor (to manage output), Li-polymer rechargeable battery, and reusable electrode patches to deliver pulses to the skin [17]. | Addresses both general menstrual pain and cramp discomfort via distinct frequency programs. Automated switching between frequencies eliminates the need for user intervention during the therapy session. Small, portable form factor make it suitable for use during daily activities. | The modality is limited in that it only applies electrical stimulation and does not apply other relief modalities such as heat or vibrational stimulation. The device depends on electrode patches, which have a finite lifespan, leading to recurring costs and potential waste. | To ensure adequate power supply/prevent brownout, the patent’s use of a controller and microprocessor can be implemented to regulate the output of heating elements and motors. The automatic cycling and timed shut off logic can be utilized to automate device operation for safety. To improve on this system, our proposed design direction will implement a bio-responsive control loop, instead of preset intervals. |
| Electric heating pad (US12011883B2) | Improves on existing heating pad/blankets by improving: -Heat distribution (reducing variable temperatures) using temperature control heat sensing [18] -Improved contact with the body by using a flexible sheet-like heating element and a flexible double layer shell [18] Improves durability by welding components together (via RF welding, ultrasonic welding, or heat bonding) as opposed to using stitching /grommets, to avoid perforations and reduce the risk of electrical shorts. [18] | Huge improvements on traditional heating pads/blankets focused on key areas like heat distribution, user comfort, and durability.[18] Heat distribution: Has temperature control with zone-sensing (positioned in areas in contact with the body) to adjust heating and keep zones closest to the body at a lower, safer temperatures than other areas not in direct contact.[18] User Comfort: It is designed with flexible materials including the sheet-like heating element and flexible shell to conforms to the body to improve contact [18] Durability: -The shell is double layered and welded together (via RF welding, ultrasonic welding, or heat bonding) as opposed to using adhesive seals/ stitches//grommets, which can make the pad less flexible and pierce insulating materials, increasing the risk of electrical shorts or burns [18] | Large/Bulky: -As it is currently designed this is not an easily portable device with large components, making the overall size roughly that of a very thin sheet cushion.[18] Expensive: Likely a very costly product given the cost of flexible heating elements [19] and the premium materials used in the design, as well as the costly technologies required to weld the components together.[18] | This design has great features which can be translated into our product like the flexible heating element, and premium flexible fabrics which are fire resistant and evenly distribute heat. Using zone sensing temperature control and placing those sensors closest to the body could also improve both heat distribution and safety. Furthermore, welding our seams shut as opposed to sewing them would enhance durability and safety. Although implementing these features/improvements would be costly, they could significantly improve the user comfort and safety of our final production-ready device. |

# Potential Design Directions

The proposed design concept centers on a multi-modal therapeutic system comprised three primary technical components: a flexible wearable belt, and bio-responsive mobile application, and a cloud-synchronized data ecosystem. The device functions through a closed-loop feedback mechanism where the user initiates therapy by logging and ranking their pain after it’s onset within the app. The user input trigger activates the wearable belt’s physical relief modalities, consisting of heat and vibrational massage therapy. The app then guides the user through targeted diaphragmatic breathing and physical exercises designed to relieve pelvic pain. The system cycle is terminated by prompting the user to re-rank their pain; if the pain has subsided, the device automatically shuts off, whereas persistent pain triggers a repeat of the therapeutic cycle.

Component I:  Hardware — Integrated Wearable Interface

The first design component focuses on the hardware integrated wearable interface, which consists of a low-profile, flexible belt designed for long-term wearability and discrete pain intervention. To address the user need for alternatives to traditional pharmaceutical pain management strategies, the belt utilizes embedded heating elements and vibration motors to provide dual modality physical relief. Previous research indicates these interventions as suitable for relieving chronic pelvic pain and dysmenorrhea[9]. Given the unpredictable nature of chronic endometriosis-related pain, there is a need for a device capable of functioning within various environmental contexts. While existing solutions are often bulky, non-discrete, or reliant on connections to power outlets, this mobile, discrete belt allows for portable pain management, which is a highly worthwhile direction to pursue for the specific user population.

Component II: Software – Relaxation Exercises via App

In addition to the hardware components that address CPP pain, the second design component also targets pain, but from a psychological perspective. Relaxation and biofeedback techniques can serve as alternatives to pharmacological interventions for managing CPP pain. They have been shown to reduce pelvic floor pain and promote muscle relaxation by supporting regulation of the autonomic nervous system response [20]. Studies have also demonstrated that relaxation exercises incorporating diaphragmatic breathing are equally effective when delivered in person or remotely, which supports the development of an app to guide users through these sessions  [13]. By creating a multimodal approach that combines a wearable device with physical pain-relief elements and techniques that target the autonomic nervous system, the overall impact has the potential to be significantly greater.

Component III: Bio-Responsive Control Loop

The second component focuses on the logic-driven software that governs the device’s operation based on real-time user input. Software design directions address user needs in both safety and psychological support in pain management. This component prioritizes patient safety by addressing the risk of “toasted skin syndrome” or cutaneous injury associated with prolonged heating pad use [11]. By employing an automated shut-off mechanism that only re-activates if the user confirms persistent pain after the exercise phase, the design ensures that heat therapy is active only when clinically necessary. Additionally, the integration of in-app guided exercises in both diaphragmatic breathing and stretches indicated in relieving pelvic pain offers both physical and psychological relief to users. Both diaphragmatic breathing and physical exercise including yoga have proven to provide pelvic pain relief  [21], [22], [23]. Furthermore, meditative breathing reduces pain perception by activating the parasympathetic nervous system, reducing muscle tension, and improving oxygen flow [21]. By integrating these guided exercises, the software integration offers psychological support to users by addressing the fear-avoidance cycle common amongst patients with chronic pain. A bio-responsive design direction is a significant improvement over standard timers, providing a safer and more empathetic user experience.

Component IV: Data Collection and Community Ecosystem

Finally, this component involves a cloud-synchronized platform for longitudinal tracking and patient experience validation. Tracking symptoms for dysmenorrhea is an effective way to advocate for oneself as it provides evidence for care and allows the patient to understand the full scope of their symptoms for better interventions. The data logged by users—including pain frequency, intensity, and relief patterns—can be exported for healthcare providers to facilitate clinical decision-making and remote monitoring. By quantifying subjective pain experiences, both clinicians and patients gain access to insights into chronic pain associated with endometriosis. Additionally, a secure in-app community forum allows users to share experiences, addressing the significant validation-gap often experienced by women seeking endometriosis diagnoses and support. This direction facilitates proactive interventions and validates the patient’s lived experience.

Figure 1: Design Flowchart for Device

# Technical Specifications

| Need | Specification | Tolerance & Target Value | Justification | Source |
| --- | --- | --- | --- | --- |
| Safe therapeutic heat for pelvic pain relief | The device shall maintain a skin-contact temperature between 38°C and 45°C during active therapy and shall not exceed 45°C under any condition. | 38°C – 45°C ±1°C MAX at 45°C ±1°C | Heat in this range provides therapeutic benefit while temperatures above 45°C increase burn and erythema risk[11], [24]. | User Comfort & Safety |
| Prevent thermal injury from prolonged exposure | The device shall automatically terminate heat delivery after ≤20 minutes per session and shall not allow more than 3 consecutive sessions without a 30-minute cooldown period. | ≤20 min per session; MAX 3 cycles | Limits cumulative thermal exposure and reduces risk of chronic heat injury (“toasted skin syndrome”) [11] | User Safety |
| Ensure discreet wearability | The wearable belt shall have a maximum thickness of 50 mm and a total weight ≤500 g. | ≤50 mm; ≤500 g | Reduces social visibility and improves long-term adherence. | User Comfort |
| Log pain onset and levels | The app should provide an interface with a likert scale from 1-10 to prompt users to enter pain level when the device is activated. | Scale should appear 1s after device is activated. Tolerance: ≤2s | Long wait times reduce app usability and enjoyability ratings. | User Design and Comfort Considerations |
| Data is stored for future access by clinical providers | The app should securely store all data collected from the user. This data includes initial pain, whether the user selected guided exercises, length of session, and final pain score. | 100% of data is maintained, tolerance of 5%. | Data can be safely stored to protect user privacy and support additional research, or documentation that clinicians require | User privacy, Healthcare providers, Researchers |
| Enable easy wearability and removal of the device | ≥80% of untrained users shall be able to put on and remove the device in ≤30 seconds without assistance or tools. | ≥80% success in ≤30 seconds | Reduces frustration and supports use during acute pain episodes. | User comfort and access |
| Ensure reliable symptom logging | ≥90% of pain log entries shall transmit successfully to the paired mobile device within 15 seconds of submission. | ≥90% transmission success within 15 seconds | Reliable data is required for clinician monitoring and longitudinal analysis in addition to quick relief for users. | Research, User Pain Management, App Developers |
| Safe and reliable portable power delivery | The device shall draw no more than 5.0 A continuous current during maximum-output operation (combined heat and vibration) and shall automatically disable heat and vibration when pack voltage falls to ≤6.0 V. | I_max ≤ 5.0 A (continuous); V_cutoff = 6.0 V ±0.2 V | Limiting current prevents battery overheating and stress. A 6.0 V cutoff (≈3.0 V per cell) prevents over-discharge, improving battery safety and longevity. | User Safety, Battery Pack Specification |

# Budget

Materials:

| Component | Product Name/Number | Purpose | Source | Unit Price (CAD) | Required # | Total Estimated Cost |
| --- | --- | --- | --- | --- | --- | --- |
| Microcontroller | ESP32 development board (with headers)à ideal b/c built in Bluetooth, small form factor, low cost (ABX00083) | Control system/brain: Runs app commands, reads sensors, controls heater and vibration output. | Digikey | $27.89 | 1 | $27.89 |
| Heating element | SparkFun 11289 | Small, flexible heating element. Provide low level heat therapy. | Digikey | $10.10 | 1 | $10.10 |
| Temperature sensor | Adafruit MCP9808 breakout | Measures heater temp, required for monitoring for safety limits. | Digikey | $8.22 | 1 | $8.22 |
| Vibration motor | Vybronics VG1030001XH LRA motor | Generate the oscillations/vibratory force. | Digikey | $4.48 | 3 | $13.44 |
| Motor driver | Adafruit DRV2605L | Required to drive the vibration motors. -supplies correct voltage/current -Enables controlled vibration (including on/off and intensity) | Digikey | $11.48 | 1 | $11.48 |
| Power supply | Powerbank 10000 mAh | Supplies power min requirements: 5 V output, >= 1.5 A output, rechargeable, USB-C included, small size. | Supplied by group member (Hanna) | N/A | 1 | N/A |
| Power switch | MOSFETs: IRLZ44NPBF | Acts as electronic switch (controlled by ESP32). Allows PWM control of heater. Enables temperature modulation. Turns heater and motor ON/OFF. | Digikey | $2.74 | 2 | $5.48 |
| Connectors – male headers | Connector header, through hole, 10 position, solder termination | ESP32 to perfboard connection, MOSFET, sensor, motor connections. | Digikey | $3.03 | 3 | $9.09 |
| Wiring – female to female | TW-E012-000 | Allow removable electronics module via wiring. Simplify repairs. | Digikey | $9.69 | 1 | $9.69 |
| Wiring – heat tolerant silicone wires | 22 AWG silicone wire | For higher current paths (ie heater, motor, battery). | Digikey (red) Digikey (black) | Per 1 ft: Red $1.13 Black $1.59 | Red = 3 ft Black = 3 ft | Red = $3.39 Black = $4.77 TOTAL = $8.16 |
| Perfboard/protoboard | DKS-SOLDERBREAD-02 | Hold/mount electronics together. | Digikey | $5.04 | 1 | $5.04 |
| Insulation barrier | Heat resistant them fleece-insultating liner | Provide heat resistant barrier between electronics and heat component. | Walmart | $8.97 | 1 | $8.97 |
| Internal fabric | Neoprene fabric | Waterproof (nonporous), fire resistant fabric to hold heating element | Amazon | $16.4 (45x 20cm) | 1 | $16.40 |
| Liner fabric (skin to device contact point) | Tightly woven Cotton Flannel | -Soft on skin -Washable -Fire resistant | Amazon | $19.99/yard | 1 | $19.99 |
| Belt/waistband fabric | Adjustable elastic strap | belt/waistband portion | Amazon | $9.99 | 1 | $9.99 |
| Fastener | Side release buckle | For belt attachment point | Supplied by group member (Hanna) | N/A | 1 | N/A |
| TOTAL ESTIMATED COSTS |  |  |  |  |  | $163.94 |

# Timeline (Gantt Chart attachment)

# Project Risk Analysis

| Risk Category | Risk Description | Likelihood | Impact | Mitigation Plan |
| --- | --- | --- | --- | --- |
| Technical (Sensor integrity and thermal control) | Failure, drift, or latency in temperature sensor telemetry (Adafruit MCP9808) may result in inaccurate thermal feedback, potentially allowing the heating element to exceed the enforced safety threshold (42°C). | Medium | High | Implement real-time temperature monitoring (1–5 s intervals), enforce a firmware-level hard cutoff at 42°C independent of app control, validate TMP117 sensor calibration prior to integration, and trigger an automatic transition to ERROR → IDLE state upon threshold violation. |
| Technical | Vibration motor intensity control may exceed safe or comfortable levels due to firmware or app communication errors. | Medium | Medium- High | Set firmware-based vibration intensity cap (0–3 levels), validate motor output during testing, and include automatic shutdown for overload detection. |
| Technical | Power supply may be inadequate, or incompatible with selected components, including the heat and vibration elements, and selected wire. | Medium | High | For each component, ensure calculations for voltage, current, resistance, etc. are completed prior to device assembly to ensure adequate power budgeting. |
| Technical | Electronics module may fail due to exposure to high temperatures from the heat element. | Low-medium | High | Implement insulating barriers between heat and electronics components: -Thermal barrier (aluminum silica insulation, neoprene fabric) -Electrical insulation (exposed wiring/headers insulated with heat resistant shrink tubing) -3D printed enclosure for electronics (Arduino, MOSFET, DRV260L board) |
| Prototyping | Integration between mobile app commands and device firmware (heat + vibration control logic) may not operate as expected. | Medium | High | Individually test firmware, app interface, and command-response protocol before full system integration |
| Safety | Device may fail to auto-shutdown during overheating, sensor failure, or prolonged session duration. | Low-Medium | High | -Overheat detection -Sensor failure detection -Max session duration (30 minutes) -User STOP override |
| User related | The belt style prototype is uncomfortable or difficult to wear | Medium | Medium | Refine the design with soft materials and adjustable straps, and collect informal feedback from users |
| Schedule | Delays due to shipping or late delivery of components | Medium | Medium | Order components early and identify local suppliers or lab alternatives |
| Team | Uneven workload distribution or team member availability issues | Low-Medium | Medium | Clearly assign roles early and cross-train team members on key tasks |
| Time | Insufficient time to fully validate prototype performance | Medium | High | Assign specific responsibilities and focus on building proof-of-concept features before optimizing. |

References:

[1]	H. Canada, “Government of Canada strengthens access to sexual and reproductive services for people living with Endometriosis.” Accessed: Feb. 02, 2026. [Online]. Available: https://www.canada.ca/en/health-canada/news/2023/09/government-of-canada-strengthens-access-to-sexual-and-reproductive-services-for-people-living-with-endometriosis.html

[2]	C. Allaire, M. A. Bedaiwy, and P. J. Yong, “Diagnosis and management of endometriosis,” CMAJ, vol. 195, no. 10, pp. E363–E371, Mar. 2023, doi: 10.1503/cmaj.220637.

[3]	B. Elbasueny, M. Geerts, E. C. Yang, C. Allaire, P. J. Yong, and M. A. Bedaiwy, “Medical treatments of endometriosis: a review,” Reprod. Dev. Med., vol. 7, no. 3, p. 166, Sep. 2023, doi: 10.1097/RD9.0000000000000053.

[4]	A. A. De Graaff et al., “The significant effect of endometriosis on physical, mental and social wellbeing: results from an international cross-sectional survey,” Hum. Reprod., vol. 28, no. 10, pp. 2677–2685, Oct. 2013, doi: 10.1093/humrep/det284.

[5]	I. Chen et al., “Hospital-associated Costs of Chronic Pelvic Pain in Canada: A Population-based Descriptive Study,” J. Obstet. Gynaecol. Can., vol. 39, no. 3, pp. 174–180, Mar. 2017, doi: 10.1016/j.jogc.2016.12.008.

[6]	P. Vercellini, C. Buffo, P. Viganò, and E. Somigliana, “Update on Medical Treatment of Endometriosis: New Drugs or New Therapeutic Approaches?,” Gynecol. Obstet. Invest., vol. 90, no. 6, pp. 535–559, Dec. 2024, doi: 10.1159/000542947.

[7]	L. C. Giudice, “CLINICAL PRACTICE,” N. Engl. J. Med., vol. 362, no. 25, pp. 2389–2398, Jun. 2010, doi: 10.1056/NEJMcp1000274.

[8]	K. P. Korekar, P. O. B. Jamal, and K. A. Borisovna, “The PHYSIOLOGY OF MENSTRUAL PAIN AND THE EFFECTIVENESS OF TECHNOLOGY-BASED PAIN MANAGEMENT TECHNIQUES,” Spec. J. Med. Acad. Life Sci., vol. 1, no. 5, Jun. 2023, doi: 10.58676/sjmas.v1i5.32.

[9]	M. Hoseini, S. Gharahtapeh, and A. Jahazi, “EFFECT OF VIBRATION AND HEAT COMBINATION ON PRIMARY DYSMENORRHEA,” Bali Med. J., vol. 4, Feb. 2015, doi: 10.15562/bmj.v4i1.99.

[10]	N. Sharma, K. Rekha, and J. K. Srinivasan, “Efficacy of transcutaneous electrical nerve stimulation in the treatment of chronic pelvic pain,” J. -Life Health, vol. 8, no. 1, pp. 36–39, 2017, doi: 10.4103/jmh.JMH_60_16.

[11]	F. Frincu, L. Scurtu, A. Bauşic, A. Petca, and C. Mehedintu, “Erythema ab igne—A Potential Cutaneous Marker of Chronic Heat Use in Patients with Endometriosis: A Narrative Literature Review and a Case Report,” Life, vol. 15, p. 1533, Sep. 2025, doi: 10.3390/life15101533.

[12]	R. J. Gatchel, R. Neblett, N. Kishino, and C. T. Ray, “Fear-Avoidance Beliefs and Chronic Pain,” J. Orthop. Sports Phys. Ther., vol. 46, no. 2, pp. 38–43, Feb. 2016, doi: 10.2519/jospt.2016.0601.

[13]	E. Peper, R. Harvey, S. Chen, and N. Heinz, “Practicing Diaphragmatic Breathing Reduces Menstrual Symptoms Both During In-Person and Synchronous Online Teaching,” Appl. Psychophysiol. Biofeedback, Oct. 2025, doi: 10.1007/s10484-025-09745-7.

[14]	Л. Э. А. ДЕВАГЕНАРЕ, “System and device of thermal treatment with protocol based on biological feedback (RU2565099C2),” RU2565099C2

[15]	M. Pisharodi and D. Rabindran, “Cold/heat-assisted distributed wave vibration therapy,” US10123937B2, Nov. 13, 2018 Accessed: Jan. 30, 2026. [Online]. Available: https://patents.google.com/patent/US10123937B2/en

[16]	W. J. Lee, H. J. Kim, C. S. Han, and S. W. Kang, “Portable combined stimulation device for alleviating menstrual pain,” US9421123B2, Aug. 23, 2016 Accessed: Feb. 22, 2026. [Online]. Available: https://patents.google.com/patent/US9421123B2/en

[17]	R. Chopra and P. Chopra, “Wearable device to provide simultaneous relief from menstrual pain and cramp discomfort using dual frequency,” WO2023053148A1

[18]	S. D. Augustine et al., “Electric heating pad,” US12011883B2, Jun. 18, 2024 Accessed: Feb. 04, 2026. [Online]. Available: https://patents.google.com/patent/US12011883B2/en?q=(heating+pad)&oq=heating+pad+

[19]	“PLM-203/10-P,” DigiKey Electronics. Accessed: Feb. 23, 2026. [Online]. Available: https://www.digikey.ca/en/products/detail/omega/PLM-203-10-P/25640143

[20]	A. Kori, A. Singh, Y. S. Caroicar, P. E. Thomas, and A. Chandra, “The Intersection of Psychology and Gynecology: A Comprehensive Review of Pain Management Strategies in Women’s Health,” Cureus, vol. 17, no. 6, p. e86344, doi: 10.7759/cureus.86344.

[21]	J. Golec, S. Gamrot, M. Michalik, I. Sulowska-Daszyk, M. Nowak, and J. Balicka-Bom, “Effects of Diaphragmatic Therapy on Pelvic Floor Muscle Activity, Stress Levels, and Sexual Life Satisfaction in Polish Women,” Appl. Sci., vol. 15, no. 22, Nov. 2025, doi: 10.3390/app152212055.

[22]	R. Saxena, M. Gupta, N. Shankar, S. Jain, and A. Saxena, “Effects of yogic intervention on pain scores and quality of life in females with chronic pelvic pain,” Int. J. Yoga, vol. 10, no. 1, pp. 9–15, 2017, doi: 10.4103/0973-6131.186155.

[23]	E. Awad, H. A. H. Ahmed, A. Yousef, and R. Abbas, “Efficacy of exercise on pelvic pain and posture associated with endometriosis: within subject design,” J. Phys. Ther. Sci., vol. 29, no. 12, pp. 2112–2115, Dec. 2017, doi: 10.1589/jpts.29.2112.

[24]	M. D. Akin, K. W. Weingand, D. A. Hengehold, M. B. Goodale, R. T. Hinkle, and R. P. Smith, “Continuous Low-Level Topical Heat in the Treatment of Dysmenorrhea,” Obstet. Gynecol., vol. 97, no. 3, p. 343, Mar. 2001.